User's Manual

250 Mhz Fiber Optic Timing Board  ARC-22

- Introduction

Guide to the DSP software file names:

There are several input and output files tfor supporting the operation of the timing board:

timhdr.asm		A simple file consisting entirely of equate commands that

are used throughout the remaining files.

timboot.asm		This file generates the EEPROM code that is loaded into the DSP 			when the system boots. It contains DSP configuration and 				initialization procedures, a command interpreter and execution 			code for the following commands: test data link, read memory, 			write memory, load application and stop idle clocking.

tim.asm		This has the code for reading out the array, clocking the array in 			the idle state, readout related parameters, and a table of all 				application commands.

timCCDmisc.asm	Targeted towards CCD applications this file has a medly of 			routines not directly involved with array readout, such as exposure 			control, power on/off and parameter configuration control.

Generic.waveforms	Here are the definitions of the waveform tables that are read by the  			timing board DSP and written to the controller hardware to control 			the clock driver and DC bias voltages, and all operations of the 			readout. It is very specific to the particular array under control

timrom.asm		A simple file called by the "timrom" script that just includes the 			files "timhdr.asm" and "tim.asm" for generating EEPROM code.

tim			A simple script for generating a file that is downloaded to the 			timing board named *.lod that takes all the files listed above as 			input source code.

timrom			A simple script for generating a file named tim.s that is suitable 			for programming the timing board EEPROM in an external 			programmer and the file named 'tim.rom' that can be written to 			the EEPROM by the DSP over the fiber optic link with the 			program TimRomBurn.

tim.lod			The output file from the "tim" script, this file can be downloaded 			over the optical fiber link into the timing board DSP memory for 			execution.

tim.rom		Similar in format to the "tim.lod" file this is produced as an output 			of the "timrom" script and is suitable for writing to the timing 			board EEPROM with the "TimRomBurn program.

tim.s			Produced as an output of the "timrom" script this file is in 				Motorola S-record format suitable for being written to the timing 			board EEPROM with external programmers.

Command Protocol:

The command and reply protocol is the same as all previous boards, with the number 2 assigned to this board. Command strings to the timing board must be in the following format, where $ designates hex and all words of the command string are 24-bits in length:

$ss02nn	'command'	arg1  	arg2 	arg3  	arg4	arg5

ss		source number of the entity issuing the command

nn		total number of words in the command, 7 >=  nn  >= 2

command 	three character ASCII command, always upper case

arg# 		optional 24-bit arguments

Command processing is all done in code located in the "timboot.asm" file, which is loaded into the DSP memory when the DSP boots after power-up. Configuration Control:

There are "Controller Configuration" bits set up in the timing board application command that can be read by the host computer to determine the capabilities of the controller hardware and software. These bits are encoded into a 24-bit word read by the computer with the 'RCC' = 'Read Controller Configuration' command. The configuration control bits are defined with equate directives in the "timhdr.asm" file, and below is shown an example of the usage with the video processor.

In the "timhdr.asm" file there are several lines that define the configuration bits:

BIT #'s		FUNCTION

2,1,0		Video Processor

000	CCD Rev. 3

001	CCD Gen I

010	IR Rev. 4

011	IR Coadder

100	CCD Rev. 5, Differential input

CCDVIDREV3B	EQU	$000000		; CCD Video Processor Rev. 3

VIDGENI		EQU	$000001		; CCD Video Processor Gen I

IRREV4		EQU	$000002		; IR Video Processor Rev. 4

COADDER		EQU	$000003		; IR Coadder

CCDVIDREV5	EQU	$000010		; CCD Video Processor Rev. 5

In the file "tim.asm" for the generic CCD controller these bits are used to define "CC":

CC	EQU	CCDVIDREV3B+TIMREV5+SHUTTER_CC

CONFIG	DC	CC		; Controller configuration

and in the file "tim.asm" the 'RCC' command is entered in the command list,

DC	'RCC',READ_CONTROLLER_CONFIGURATION

while in the file "timCCDmisc.asm" the 'RCC' command is finally defined:

; Let the host computer read the controller configuration

READ_CONTROLLER_CONFIGURATION

MOVE	Y:<CONFIG,X0		; Just transmit the configuration

JMP	<FINISH1

Software Status Bits:

Also in the file "timhdr.asm" are defined bits for the STATUS word located at X: memory location 0 of the timing board DSP. These are mainly used to control execution of the timing board code with bit conditional jump instructions (JCLR and JSET). Some of the bits are set by the host computer by reading the X:STATUS word, changing the desired bit, and writing the word back to the controller. Other bits are controlled by the DSP to indicate execution status. Such things as shutter mode, MPP operation, CCD readout architecture are controlled with these bits.

Power ON and OFF:

The file "timCCDmisc.asm" includes commands for powering the power control board on and off, which causes the analog supply voltages (+/-6.5V, +/-16.5V and +36V) to switch these voltages onto or off of the backplane. These commands ('PON' and 'POF') are also resident in the utility board, so the host computer can issue them to either board and have them execute properly. This capability is included to protect the image sensors from damaging voltage combinations or sudden voltage spikes, as the power control board turns on the power in sequence with writing voltage values to the DACs and closing output protection switches on the analog boards. Also, the power control board turns on the voltages with a timescale of about 10 milliseconds. Note that the +5V digital power is not switched but instead is turned on at all times to all the boards in the controller.

There is a routine named CLEAR_SWITCHES that writes zeroes to all the analog switches on the video processors and clock driver boards. This is needed to minimize power dissipation in these parts that occurs when their +5V digital control lines are powered but their analog supplies are not. This same routine is executed in the "timboot.asm" code when the DSP first boots up. The CLEAR_SWITCHES is also called on power off.

The power on procedure consists of the following:

1. Write zeroes to all the analog switches and DACs.

2. Write zeroes to the LVEN (Low voltage enable, for the +/-6.5V and +16.5V 		supplies), and then to HVEN for the +36V supply to turn them on.

3. Check the power control board signal POWEROK to ensure that the power 		was indeed turned on correctly. If not, report an error.

4. Close the output analog swtiches that connect the clock driver and DC bias 		supply output circuits to the sensor.

5. Write all DAC values from the "DACS" table to the DACs. The "DACS" table 		is located in the "Generic.waveforms" file, and delay.

Shutter control:

Shutters can be controlled from the timing or utility boards. Shutter control with the utility board is accomplised with an open collector circuit (74F07 and a pull up resistor supplied near the shutter driver circuit by the user) driving the SHUTTER line on the front of the utility board available on the small printed circuit board that plus into its DIN-96 connector and is wired to a small circular 6-pin HIROSE connector mounted next to the fiber optic connectors on the controller housing. Alternatively, shutters can be controlled with the timing board, wherein the shutter control signal, also named SHUTTER, is routed from U12 pin #15 to the backplane pin A30. From here it is routed through the power control board to the small power supply to a shutter driver circuit that translates this TTL level signal to a +24-volt burst followed by a +5 volt holding signal to open solenoid-driven shutters. This shutter driver signal is routed from the small power supply box back to the power control board were it is wired to the small circular HIROSE connector mounted next to the fiber optics connectors on the controller housing. Finally, the TTL-level signal SHUTTER from the timing board is also available directly on the power control board (only for Rev. 7 power control boards), which is also routed to the same HIROSE circular connector. The HIROSE connector is wired as following:

pin #1 - White wire to utility board front connector, ground

pin #2 - Grey wire to utility board front connector, open collector shutter

pin #4 - White wire to power control board J6, +24 volt shutter driver pulse

pin #5 - Grey wire to power control board J6, +24 volt shutter driver pulse

pin #3 - Blue wire to power control board J8 pin #2, ground

pin #6 - Purple wire to power control board J8 pin #1, TTL SHUTTER

The shutter control software is located in the file "timCCDmisc.asm". The shutter is opened by writing a zero to bit #4 of the LATCH (which is connected to U12) and a one to close the shutter, taking care not to disturb the other bits. The shutter can be controlled directly by the host computer with the 'OSH' and 'CSH' commands, or from DSP software by the subroutines 'OSHUT' and 'CSHUT'. The shutter is opened at the beginning of an exposure and closed at the end of it if the SHUT bit #11 of the X:STATUS word is set. Code in the "tim.asm" file delays readout by Y:SHDEL milliseconds after the shutter is first instructed to close to allow it to close completely before charge is moved on the CCD to prevent image smearing.

Exposure Sequence:

The sequence of commands and software execution from the beginning of an exposure to its completion will be described here. The sequence will be described assuming that the timing board is controlling the exposure. Alternatively the utility board could be used, and its operation is described in the utility board user's manual.

1. The host computer sets the desired exposure time by issuing the 'SET' = Set exposure time command to the timing board with an argument of milliseconds.

0x203		'SET'		Number of milliseconds

2. The timing board executes the SET_EXPOSURE_TIME located in "timCCDmisc.asm" routine by simply storing the value in the EXPOSURE_TIME constant.

3. The host computer sets the shutter status to either open the shutter or not during the exposure by reading the X:STATUS value, setting or clearing bit #11, and writing the value back to the same location.

4. The host computer issues the 'SEX' = start exposure command to the timing board.

0x202		'SEX'		Start exposure

5. Upon receiving the 'SEX' command the START_EXPOSURE routine located in "timCCDmisc.asm" issues a 'IIA' = initialize image address command to the PCI board. The PCI board will set the PCI image address equal to the base PCI address and the number of pixels transferred to zero.

6. Continuing in the START_EXPOSURE routine located in "timCCDmisc.asm", the CCD is cleared with the CLR_CCD routine, keeping care to periodically check for incoming commands so the abort readout or abort exposure command can be executed promptly.

7. The variable X:<IDL_ADR is written with the address #<TST_RCV so that when the timing board processes normal comamnds after the exposure begins that the timing board will not be executing the IDLE routine in the "tim.asm" file that exercises the parallel and serial clocks.

8. The shutter is opened if bit #11 = SHUT bit is set.

9. The EXPOSE routine in the file "timCCDmisc.asm" is entered, which loads the EXPOSURE_TIME into the DSP timer register, enables the timer, and monitors the timer status for the end of the timer countdown. During the countdown incoming commands are checked for in case an abort expsoure command is issued, and the serial clocks are exercised, but not the parallel clocks.

10. When the countdown is complete the START_EXPOSURE routine is re-entered, causing the shutter to be closed if the SHUT bit of X:STATUS has been set, and causing execution to jump to the RDCCD routine in the file "tim.asm".

11. At the start of the RDCCD routine in "tim.asm" a bit is set to indicate that readout has begun and the PCI_READ_IMAGE routine in "timCCDmisc.asm" is executed. This routine sends the 'RDA' command with arguments containing the total number of pixels to be read in the image to the PCI board that sets up a counter that awaits these total number of pixels until it signals to the host computer that image readout has completed. Notice that with this revision 1.7 of the software that the host computer does not write any image size information to the PCI board, but instead writes this to the timing board that then writes it to the PCI board.

12. Continuing in the RDCCD routine in "tim.asm", a SYNTHETIC_IMAGE routine in "timCCDmisc.asm" will be executed if the TST_IMG bit #10 of X:STATUS is set, which will cause an image consisting of incremented pixel counts to be transmitted to the PCI board. If a noirmal readout is desired then there is a delay of Y:SHDEL milliseconds to allow the shutter to close, a clearing out of the charge in the serial shift register, and then a simple readout of Y:NPR rows and Y:NCOLS in each row. The PARALLEL and SERIAL_READ routines that are called to exercise the clocks are located in the "Generic.waveforms" file.

13. At the end of the readout the X:<IDL_ADR address is filled with the address of the IDLE routine if the #IDLMODE bit of X:STATUS has been set by the host computer, the reading out in progress bit is cleared and execution jumps to command processing.

14. Meanwhile the PCI board has been reading image data from the fiber optics data link and writing them to sequential PCI addresses, counting the number of pixels transferred. When the desired number has been transferred the host flags are changed from reading out to done status, and normal command execution continues. An interrupt may be generated at this point if desired.

Other exposure comamnds:

Several additional exposure commands are implemented in "timCCDmisc.asm", as follows:

'PEX' = Pause exposure. Simply close the shutter if needed and disable the 		exposure timer.

'REX' = Resume exposure. Simply open the shutter if needed and enable the 		exposure timer.

'RET' = Read exposure time. Read the current contents of the exposure, in 		milliseconds remaining until the end of the exposure.

'SET' = Set exposure time. The argument is the number of milliseconds desired 		in the exposure. This command need only be issued if there is a change 		from the previous exposure time.

'AEX' = Abort exposure. This will terminate the current exposure by closing the 		shutter, disabling the timer and sending a 'DON' reply to the PCI board.

'ABR' = Abort readout. Similar to abort readout, this will terminate the readout 		currently in progress

'CRD' = Continue readout. This will just continue the readout from where it was 		left off by the abort readout command.

Support routines:

There are several support routines included in the "timCCDmisc.asm" file intended to make setup and operation of the hardware somewhat easier. There are as follows:

1. 'SGN' = Set video processor gain. Gains of x1, x2, x4.75 and x9.5 are specified by the first command, and the speed of the integrator is set by the second argument, where 0 is appropriate for slow system (introducing 4.9 pf capacitance in the integrator, causing a low gain) and 1 is entered for a fast system (introducing a 1.0 pf capacitance in the integrator). The syntax is as follows:

0x204		'SGN'		{1, 2, 5 or 10}	{0 or 1}

Gain		  Speed

2. 'SDC' = Set DC coupling diagnostic mode. There is a DC-coupling mode on Rev. 3 video processors that needs to have an input offset voltage set properly. The output of the preamp can be routed to the A/D converter directly so this voltage can be measured easily as an aid to setting the input offset voltage. An argument of 1 is issued with the 'SGN' command to cause this to happen.

3. 'SBN' = Set bias number is useful for updating the digital value in any of the DACs in the video processors or clock driver boards. The syntax is as follows:

0x206	   'SBN'     #board     #DAC     {'VID' or 'CLK'}     12-bit value

#board = jumper setting of the board

#DAC = number of the DAC on the selected board

select for updating video processor or clock driver boards

enter the 12-bit value to be written to the DAC

4. 'SMX' = Set multiplexer number on the clock driver board will connect any two clock driver signals to SMB connectors on the front of the board, useful for viewing the signals on an oscilloscope for verifying timing and signal overlaps. The command syntax is:

0x204	   'SMX'     #board     #MUX1     #MUX2

#board = jumper setting of the board

#MUX1 = number from 0 to 23 of the clock signal to be mapped to the output 		MUX connector #1.

#MUX2 = number from 0 to 23 of the clock signal to be mapped to the output 		MUX connector #2.

5. 'CSW' = Clear analog switches to zero, to minimize their power dissipation in the 	absence of analog power.

Waveform timing:

The generic waveform timing file that is included reads out the CCD at the maximum rate allowed by the video processor, 1.0 microseconds per pixel. This rate is determined by the SERIAL_READ waveform table, where the following notes apply:

1. Each line of the waveform table takes 40 microseconds to execute.

2. The first eight bits #23-16 of the waveform word are delay bits, as follows:

If bit #23 = 1 then bits 22-16 specify the number of 640 nanosec cycles 			that bits 15-0 will be held at before the next waveform word is read.

If bit #23 = 0 then bits 22-16 specify the number of 40 nanosec delay cycles that 		bits 15-0 will be held at before the next waveform word is read.

3. The bits #15-12 select the board number to which the bits #11-0 are written to. If 	these bits match the four jumper settings of the "SWITCH" bits on either the 	clock driver or video processor boards, then the selected board will latch bits 	#11-0 into its internal registers.

4. If the bits #15-12 are all set (= 0xF) then a special selection is made of writing the 	bits #11-0 to the circuitry on the timing board that exercises the serial 	transmission of A/D video data over the fiber optics data link (SXMIT):

bits #5-0 select the starting number of the A/D converter to be 				transmitted in a sequence.

bits #10-6 select the ending number of the A/D converter to be 				transmitted in a sequence.

5. The clocking bits are bits #6-0 for the video processor and bits #11-0 for the clock 	driver. When the bit is set the clock is driven high, and low when cleared.

6. This example refers to the SERIAL_READ code from the Generic.waveform file.

SERIAL_READ

DC	END_SERIAL_READ-SERIAL_READ-1

DC	CLK2+S_DLY+SW+RG+00+S2+S3+00+S5+00	; L2 = H

DC	CLK2+00000+SW+00+00+S2+S3+00+S5+00	; RG = L

DC	VIDEO+%1110100

DC	$00F0C0					; Transmit A/D data to host

DC	VIDEO+$000000+%1110111		; Stop resetting integrator

DC	VIDEO+$050000+%0000111		; Integrate - delta

DC	CLK2+$010000+00+00+S1+00+00+S4+00+S6

DC	VIDEO+$010000+%0011011		; Stop Integrate

DC	VIDEO+$070000+%0001011		; Integrate

DC	VIDEO+$000000+%0011011		; Stop, A/D is sampling

END_SERIAL_READ

a. The first entry in a waveform table needs to be the number of words in the table, 	minus one. Rev. 4 timing boards (50 Mhz fiber optic speed) required a minus two 	here. The simple construction below is recommended. Each line defines one word 	of DSP memory, using the DC = Define Constant directive to the Motorola 	assembler.

b. The board selection is made with CLK2 = $002000 for the clock driver board and VIDEO = $000000 for the video processor board. The clock driver board signal names define one bit of the clock driver's 12-bit data word to be high, so when a zero shows up instead the clock signal is low. These names generally are close to the names assigned by the CCD manufacturer. The video processor bits are listed explicitly in this example, where the "%" signifies a binary representation of the string. The video processor bits are defined as

Bit #6 = xfer => on a low-to-high transition transfer 16-bit image data from the 		A/D converter output to a latch.

Bit #5 = A/D => on a low-to-high transition start the A/D sample/hold and 		subsequent conversion.

Bit #4 = integ => when low integrate on the dual slope integrator.

Bit #3, 2 = Pol+, Pol- => with the value 01 configure the amplifier circuit in a 		non-inverting configuration, and with the value 10 configure it inverting.

Bit #1 = Dcclamp => when low clamp the input video signal to ground

Bit #0 = rst => when low reset the integrator

c. In this example there is an integration of the video signal before the dumping of charge from the summing well to the output node with a non-inverting polarity, followed by an integration with an inverting polarity for an equal amount of time after the charge dump. Notice that a trick is played here of clocking the summing well low to dump the charge before the first signal integration is complete. Normally this is not a good thing, but in this system there was a long delay between the summing well clocking and a change in the video signal, allowing some time to be saved with this sequence. It is a good idea to verify this by examining the integrate signal on the integrator analog switch (the INTEG signal) with the video signal to ensure that the video signal transition due to summing well clocking occurs well outside the two integration times when the INTEG signal is low. Similarly, the first integration should not occur too close to the last CCD charge clocking, which is why the A/D transmisstion and SXMIT commands are place right before the first integration time.

d. The two integration times should be of equal duration. In this example they are both 320 nanosec long. The first integration time consists of a delay of 5, a seconds delay of 1, for a total delay of 6 x 40 nsec = 240 nsec plus two 40 nsec instruction times, one for the instruction to begin the integration and one for the summing well clocking instruction. The second integration time is just the delay time of 7 x 40 nsec plus the instruction time of 40 nsec.

Voltage Settings:

The voltages of the DACs on the video and clock driver boards are set in the routine named SET_BIASES in the file "timCCDmisc.asm" by reading from the DACS table in Genric.waveforms. The voltages are specified in a table in units of decimal volts, as

; Define the clocking voltages

SW_HI	EQU	+3.0	; Summing well, +6, -3 for anti-backspill

SW_LO	EQU	-8.0

and then in the DACS voltage table as

DACS		DC	END_DACS-DACS-1

DC	(CLK2<<8)+(0<<14)+@CVI((SW_HI+10.0)/20.0*4095)

DC	(CLK2<<8)+(1<<14)+@CVI((SW_LO+10.0)/20.0*4095)

These two lines construct the word that is to be transmitted over the serial SSI link to the clock driver board to be written to the selected DAC, following the protocol that is described in the clock driver user's manual. The CLK2<<8 simply left shifts the CLK2 definition of 0x002000 so the 2 is in the bit position #23-20 required for serial words, and the (0<<14) shifts the DAC number 0 into the bit locations #19-12. The next       argument is to the assembler directive @CVI for constructing a 12-bit integer for the value to be written to the DAC corresponding to a high voltage of SW_HI = +3.0 volts , and a 20 volt bipolar output range. The second line above simply writes the low clock voltage to the next DAC location, so the clock driver can switch between the two voltages according to the value of the bit in the clocking waveform.